MOISTURE IN OIL SENSOR

The compact moisture in oil sensor is designed for reliable online monitoring of the moisture in transformer, engine, lubrication or hydraulic oil as well as in diesel fuel. Besides the accurate measurement of water activity (aw) and temperature (T), the EE364 calculates the absolute water content (x) of the oil in ppm.

#### Outstanding Measurement Performance

The sensor employs high end E+E humidity sensing elements which feature outstanding long term stability and high resistance to pollution.

#### Functional Design

The small size and robust stainless steel enclosure, together with the choice of process connections allow easy and space-saving installation.

#### Analogue and Digital Outputs

The aw, T and x measured data is available on two freely configurable 4 - 20 mA analogue outputs and on the RS485 interface with Modbus RTU protocol. The wide scaling range of the analogue output facilitates the EE364 implementation in existing monitoring and control systems.

#### Configurable and Adjustable

The setup and adjustment of the EE364 can be easily performed with an optional configuration adapter and the free EE-PCS Product Configuration Software.

# Features

Measurement performance

» Water activity (aw), water content (x), temperature (T)

» Oil temperature -40...100 °C (-40...212 °F)

» Suitable for transformer, lubrication and hydraulic oil, as well as for diesel fuel

Configurable and adjustable

» Service interface

» Free configuration software

Outputs

» 2 x 4 - 20 mA output, wide scaling range

» Modbus RTU

» Industrial M12x1 connector

Inspection certificate

according DIN EN 10204-3.1

Mechanical construction

» Stainless steel enclosure and filter cap

» Pressure tight up to 20 bar

» Process connection with ISO or NPT thread

» IP65 / NEMA

# Measurement of water activity aw / water content

The moisture in oil can be expressed in absolute or relative terms.

Water activity aw is a relative measure for moisture in oil. It represents the ratio between the actual amount of dissolved water and the maximum possible amount of dissolved water in the oil at a given temperature. The aw value indicates how close to saturation is the oil. aw = 0 means dry oil (no water at all), aw = 1 means fully saturated oil. Water activity does not depend on the type of oil.

Water content x is an absolute measure for the amount of water in the oil (dissolved, emulsified or separated). The water content is usually expressed in ppm or mg water / kg oil and it is independent from the oil temperature. For assessing the degree of saturation, x must be regarded together with T.

EE364 calculates x based on the measured aw and T values. The calculation is oil dependent and requires a set of oil specific parameters. E+E offers the service of determining the oil specific parametrs, see section “Accessories” below. The parameters can be set upon order or uploaded to EE364 using the EE-PCS Product Configuration Software.

# Technical data

### Water activity (aw) / water content (x)

Working range	0...1 / 0...100 000 ppm

(ppm output is valid in the range 0...100 °C (32...212 °F))

Accuracy1) at 20°C (68 °F)	±0.02 (aw = 0...0.9) ±0.03 (aw = 0.9...1)

Response time t90 at 20°C (68 °F), in still oil, typ.	10 min.

### Temperature

Maximum working range	-40...100 °C (-40...212 °F)

Accuracy1) at 20°C in oil	±0.2 °C (0.36 °F)

Resolution	0.01 °C (0.18 °F)

## Output

Analogue output (scalable)	2 x 4 - 20 mA (3-wire)	RL < 500 Ohm Resolution	2 µA

Digital interface	RS485 (EE364 = 1 unit load)

Protocol	Modbus RTU

Default settings	Baud rate 96002), parity even, 1 stop bit, Modbus address 234

## General

Power supply class III3)	10*) - 28V DC	*) 10V+0.02*RL

Power consumption	<20 mA + load current

Electrical connection	M12x1 plug 8 poles

Pressure rating	0...20 bar (0...290 psi)

Enclosure material	Stainless steel 1.4404 (AISI 316L)

Protection rating	IP65 / NEMA 4

Filter	Stainless steel

Oil temperature	-40...80 °C (-40...176 °F) / -40...100 °C (-40...212 °F)

Ambient temperature	-40...60 °C (-40...140 °F) / -40...80 °C (-40...176 °F)

Storage temperature	-40...60 °C (-40...140 °F)

Electromagnetic compatibility	EN 61326-1 EN 61326-2-3   Industrial environment

FCC Part15 Class A	ICES-003 Class A

- The accuracy statement includes the uncertainty of the factory calibration with an enhancement factor k=2 (2 x standard deviation).

The accuracy was calculated in accordance with EA-4/02 and with regard to GUM (Guide to the Expression of Uncertainty in Measurement).

- Supported baud rates: 9 600, 19 200 and 38 400; find more details about communication setting in the User Manual and the Modbus Application Note at www.epluse.com/EE364.

- USA & Canada class 2 supply required.

# Dimensions

Values in mm (inch)

92 (3.6)

34 (1.36)

G1/2“

AD27

77 (3)

140 (5.5)

Ø12 (0.5)

Ø30 (1.2)

ISO thread

51 (2)	     	89 (3.5)

1/2“NPT

AD22

74 (2.9)

NPT thread

# Ordering Guide

|  |  |  | EE364- |
| --- | --- | --- | --- |
| Hardware | Process connection | G 1/2" ISO thread 1/2" NPT thread | PA1 PA2 |
|  | Fluid temperature -40...80 °C (-40...176 °F) -40...100 °C (-40...212 °F) |  | HM1 HM2 |
|  | Filter | Stainless steel, for flow < 1 m/s (< 3.28 ft/s) Stainless steel, for flow > 1 m/s (> 3.28 ft/s) | F13 F18 |
|  | Accessories No accessories M12x1 cable socket, for self assembly |  | AC0 AC2 |
| Software Setup - Analogue Outputs | Output 1 | Water activity aw [ ] Water content x [ppm] Temperature T [°C] Temperature T [°F] | no code MA70 MA1 MA2 |
|  | Scaling 1 low 0 Value |  | no code SALValue |
|  | Scaling 1 high 1 Value |  | no code SAHValue |
|  | Output 2 | Temperature T [°C] Temperature T [°F] Water activity aw [ ] Water content x [ppm] | no code MB2 MB67 MB70 |
|  | Scaling 2 low -20 Value |  | no code SBLValue |
|  | Scaling 2 high 80 Value |  | no code SBHValue |
|  | Units (Modbus RTU) | Metric (SI) Non-metric | U1 U2 |
|  | Oil parameterization for Mineral transformer oil water content calculation Customer specific oil |  | no code PPMxxx |

Order Example

#### EE364-PA1HM1F13AC2U1

Process connection:	1/2" ISO thread

Fluid temperature:	-40...80 °C (-40...176 °F)

Filter:	Stainless steel, for flow < 1m/s

Accessory:	M12x1 socket connector, for self assembly Units:	Metric (SI)

# Accessories

(For further Information, see datasheet "Accessories")

| Determination of oil specific parameters | ppm-cal |
| --- | --- |
| Modbus configuration adapter | HA011013 |
| EE-PCS Product Configuration Software | EE-PCS |
| (free download: www.epluse.com/configurator) |  |
| M12x1 8 pole cable socket for self assembly | HA010704 |
| M12x1 8 pole shielded connection cable, socket/flying leads |  |
| 1.5 m (4.9 ft) | HA010322 |
| 5 m (9.8 ft) | HA010324 |
| 10 m (16.4 ft) | HA010325 |
| Sampling cell with shut-off function, PN40, DN25 | HA050109 |